GeGI Driver — Messaging & Integration Interface

Handover reference for integrating this driver into PRISM. It describes the

messaging system, the full topic/service surface, message definitions, and the

two ways an external system can talk to it.

1. Messaging model at a glance

There are two messaging layers:

GeGI detector  <--(A) raw TCP binary-->  C++ driver node  <--(B) ROS 1 pub/sub + services-->  consumers / PRISM

192.168.50.109:27015                     phds_gegi_driver                                       (native ROS or rosbridge)

•  (A) Detector ↔ driver — raw TCP, binary GeGI protocol. A single TCP socket

to the detector (default 192.168.50.109:27015). Variable-length event packets

stream in; single-byte commands go out (g=start, s=stop, c=clear,

i=run-info, d=detector-info, 1–8=timed presets, …). This is internal to

the C++ node — PRISM does not touch this layer.

•  (B) Driver ↔ everything else — ROS 1 (Melodic) middleware. This is the

integration surface. Standard ROS 1 transport (TCPROS: an XMLRPC master

roscore for discovery, then a direct TCP connection per topic). All data and

control are ROS topics (pub/sub) and services (request/response).

•  Non-ROS bridge: a rosbridge websocket server on port 9090 is launched

with the pipeline. Any language can subscribe/publish/call services as **JSON

over websocket** without linking ROS. The existing Windows plotting tools in

tools/ use exactly this.

Stack: ROS 1 Melodic, catkin. The detector node is C++; the processing nodes

(spectrum, activity, heatmap, recorder) are Python 2.7.

2. Node graph & data flow

| Node | Language | Consumes | Produces |
| --- | --- | --- | --- |
| phds_gegi_driver | C++ | detector TCP stream | /compton_event, /energy_deposit, /energy_deposit_singles; detector control services |
| spectrum_node | Py | /energy_deposit | /spectrum |
| activity_node | Py | /spectrum | /activity/results, /activity/total_activity, /activity/effective_source_distance |
| spherical_heatmap_node | Py | /compton_event, /activity/results | /sphere_heatmap, /source_direction(s), /source_isotopes |
| data_recorder_node | Py | all of the above | files (N42, CSV, rosbag, manifest); recording services |

3. Topics (the data surface)

| Topic | Type | Dir | Notes |
| --- | --- | --- | --- |
| /compton_event | radiation_detector_msgs/ComptonEvent | out | one msg per 2-site Compton event (imaging input) |
| /energy_deposit | std_msgs/Float64 | out | per-event energy (keV); 1-site + 2-site-summed → spectrum |
| /energy_deposit_singles | std_msgs/Float64 | out | per-event energy (keV), 1-site only |
| /spectrum | radiation_detector_msgs/Spectrum | out | histogram, republished ~4 Hz |
| /activity/results | std_msgs/String | out | JSON activity report (see §5) |
| /activity/total_activity | std_msgs/Float64 | out | summed activity (MBq) |
| /activity/effective_source_distance | std_msgs/Float64 | out | latched; plate-derived standoff (m) |
| /sphere_heatmap | sensor_msgs/PointCloud2 | out | back-projection image (xyz + per-isotope dose fields) |
| /source_direction | geometry_msgs/PoseStamped | out | strongest source direction |
| /source_directions | geometry_msgs/PoseArray | out | all detected source directions |
| /source_isotopes | std_msgs/String | out | `"Cs-137:count | Co-60:count"` |
| /activity/n_shielding_plates | std_msgs/Int32 | in | operator sets extra shielding plates |
| /detector/dead_time_percent | std_msgs/Float64 | out | optional shared dead-time (off by default) |

Note on rates: /compton_event and /energy_deposit fire once **per detector

event** and can burst to thousands/sec. Publisher and subscriber queues are sized

at 10,000 so bursts are not dropped.

4. Services (the control surface)

Detector control (provided by the C++ node, under /detector/):

| Service | Type | Purpose |
| --- | --- | --- |
| /detector/start_acquisition | phds_gegi_driver/StartAcquisition | begin free-running acquisition |
| /detector/stop_acquisition | phds_gegi_driver/StopAcquisition | stop |
| /detector/clear_data | phds_gegi_driver/ClearData | clear detector spectrum/state |
| /detector/start_timed_acquisition | phds_gegi_driver/StartTimedAcquisition | detector-timed preset run |
| /detector/get_run_info | phds_gegi_driver/GetRunInfo | real/live time, dead-time %, count rate |
| /detector/get_detector_info | phds_gegi_driver/GetDetectorInfo | serial, temperature, bias, battery |
| /detector/toggle_bias_mode | phds_gegi_driver/ToggleBiasMode | HV bias on/off |

Recording / orchestration (provided by data_recorder_node, under /data_recorder/):

| Service | Type | Purpose |
| --- | --- | --- |
| /data_recorder/start_timed_recording | phds_gegi_driver/StartTimedAcquisition | run + save all assets for N minutes |
| /data_recorder/stop_recording | std_srvs/Trigger | stop early |
| /data_recorder/clear_all | std_srvs/Trigger | clear detector + all node state |

Each processing node also offers ~clear (and the heatmap ~save_csv) as

std_srvs/Trigger.

5. Message definitions

`radiation_detector_msgs/ComptonEvent`

std_msgs/Header header

float64 energy_kev_1

geometry_msgs/Point reading_location_1   # metres, ROS REP-105 frame

float64 energy_kev_2

geometry_msgs/Point reading_location_2

float64 cone_angle

float64 cone_angle_uncertainty

`radiation_detector_msgs/Spectrum`

std_msgs/Header header

time   endTime

uint32 realTime_s        # legacy

uint32 realTime_ms       # use this

uint32 deadTime_ms

uint32 totalCount

uint32[] spectrum        # counts per energy bin

`phds_gegi_driver/GetRunInfo` (request empty)

---

bool    success

float64 real_time_sec

float64 live_time_sec

float64 dead_time_percent

float64 count_rate_hz

string  message

`phds_gegi_driver/StartTimedAcquisition`

int32 duration_minutes

---

bool   success

string message

`/activity/results` (JSON in `std_msgs/String`) — top-level object with a

timestamp, total_activity_MBq, and an isotopes array; each element carries

isotope, energy_keV, activity_MBq, sigma_activity_MBq, valid,

net_corrected, source_distance_m, total_shield_plates, shield_transmission,

method, etc. This is the easiest single feed for PRISM to consume — parse the

JSON string, no ROS message compilation needed.

6. Message packages / build dependencies

•  `radiation_detector_msgs` — external msg package (in deps/): ComptonEvent,

Spectrum, GetSpectrum, … PRISM needs these definitions to deserialize the

imaging/spectrum topics natively (not needed via rosbridge JSON).

•  `phds_gegi_driver/srv` — this package's custom services (§4).

•  Standard: std_msgs, geometry_msgs, sensor_msgs, std_srvs.

7. Integrating with PRISM — two paths

A. Native ROS 1 node (recommended if PRISM can host or reach a ROS master).

Point PRISM at the ROS_MASTER_URI, build against radiation_detector_msgs +

phds_gegi_driver, then subscribe to the topics and call the services directly.

Lowest latency, full-rate event access.

B. rosbridge websocket (recommended if PRISM is non-ROS / cross-platform).

Connect to ws://<host>:9090 and exchange JSON (op: subscribe/publish/call_service).

No ROS build dependency; works from any language. Already running in the pipeline

(gegi_full_pipeline.launch) and used by the tools/plot_live_*.py clients as

reference implementations. Caveat: JSON encoding of high-rate /compton_event

and binary PointCloud2 is heavier — for bulk event throughput prefer path A,

and use rosbridge for the digested feeds (/activity/results, /spectrum,

/source_isotopes, dose).

8. Practical notes for the integrator

•  The detector allows one TCP connection; the C++ node owns it. Do not open a

second connection to the detector — go through the ROS interface.

•  /activity/results and /source_isotopes are the highest-value digested feeds

for a supervisory system; raw /compton_event is only needed for custom

imaging.

•  /activity/effective_source_distance is latched — a late subscriber still

receives the last value.

•  Control flow for a measurement: call /data_recorder/start_timed_recording

(it orchestrates the detector run and writes all assets), or drive the detector

directly via /detector/* if PRISM manages its own acquisition lifecycle.

•  Everything is ROS 1 Melodic; there is no ROS 2 / DDS layer in this driver.